## Slide 1: Yokogawa InfraSpec NR800
FT-NIR
Analisador Multivariável
Notes: This presentation is intended to introduce the Yokogawa InfraSpec NR800 FT-NIR

## Slide 2: InfraSpec NR800
Notes: Photos of the process unit mounted on a roll-around stand, and the lab bench-top unit.
Note: The process unit makes a great lab unit when mounted on such a stand. For some customers, there are advantages to having the lab and process analyzers identical so that the lab unit can be used as a “hot spare” if necessary. Models are transferable between all NR800 models in any event.

## Slide 3: A Descoberta de Herschel (1800)
Notes: In 1800, William Herschel assembled a spectrometer in which the Sun served as the source, a prism as the dispersing element, and a thermometer as the detector. Having first measured the distribution of radiant heat across the visible region, Herschel took the unprecedented step of moving the thermometer to a position beyond the red end of the spectrum. To his surprise, the thermometer not only registered heat, but a higher temperature than he found in the visible region. This signaled the discovery of a “non-visible spectrum”, and coincidentally, the first NIR spectrometer.

## Slide 4: O que é NIR?
Forte
Fraca
Absorbância
Infra Spec NR800
2,5
0,9
 = 1 /  (cm-1)
 = c/ 
c = Vel.da Luz = 3 x1010 cm/sec
Far-IR
Middle-IR
Near-IR
100
25
2,5
0,75
Comp. de Onda λ µm
2,5
0,75
1,8
1,1
I
II
III
Combinação
C-H, N-H,
O-H, C=O
Sobretom
1o C-H, N-H
2o C-H
Sobretom
2o N-H
3o C-H
µm
Núm. de Ondas  cm-1
100
101
102
103
104
105
106
Comp. de Onda λ µm
1.000
100
10
1
0,1
0,01
0,001
Raio-X
Ultravioleta
Visível
Microondas
Infravermelho
Regiões do Espectro
Transição Elétron
Vibração
Molecular
Rotação
Molecular
Excitação
Elétron
107
10.000
10-1
100.000
Forte
Fraca
Influência da água
Luz vermelha => 0,7 µm = 7 x 10-5 cm
 = 1 / 7 x 10-5 cm = 1,4 x 104 cm-1
 = 3 x 1010 cm/s
7 x 10-5 cm
 = 0,4 x 1015 = 0,4 PHz (petaHertz)
Notes: Infra-red light is the region of electromagnetic radiation with wavelengths longer than visible light, just beyond the color red. The full infra-red region extends from about 0.75 micrometers (750 nm or 13,333 cm-1) to 100 micrometers (100,000 nm or 100 cm-1). The full IR range is divided into the Near-IR, Mid-IR, and Far-IR).
IR energy is absorbed by chemical bonds between dissimilar atoms (C – H, N – H, C=O, O – H, etc.) which causes these bonds to stretch and vibrate. The fundamental vibrations of these bonds occurs in the Far and Mid-IR regions, but absorbance peaks occur at discrete multiples (2, 3 , 4 …) of the fundamental frequencies in the NIR region, called “overtones”. In addition, the fundamental stretches cause secondary bending and harmonic vibration in nearby chemical bonds, which leads to “combination zones” of absorption in the NIR region.
Far-IR instrumentation is used by NASA to measure the chemical composition of stars. Mid-IR spectrometers are very common in laboratories, used for qualitative and quantitative analysis. Specialized equipment with water-sensitive salt windows (such as KBr) and other optical components are required for a spectrometer that operates in the Mid-IR or Far-IR. Generally, Near-IR equipment can have glass or sapphire components that are more practical and durable for process applications. Absorption in the Mid-IR tends to be very strong and in narrow bands. NIR absorption is weaker, in broader bands that tend to interact in complex ways. This weaker more complex absorption is advantageous in many process monitoring applications for very practical reasons. Also, the influence of water, which absorbs more strongly as you approach the Mid-IR, is less, making a large section of the NIR spectrum available for process measurements without interference from water.

## Slide 5: Absorção no Infravermelho
Vibração das ligações Químicas
Notes: Light energy is absorbed in the Infrared by the vibrations of chemical bonds.
The fundamental frequencies of vibration for most chemicals bonds occur in the Mid-IR. The absorbance at these frequencies is very strong, which leads to strong peaks of absorbance, which are characteristic for these bonds. Thus, Mid-IR is a powerful tool for the identification and quantification of chemicals and has been used in laboratories for many years.
The vibration of chemical bonds is not a simple harmonic motion. Due to the difference in size and mass of two dissimilar atoms bonded, weaker overtones of bond vibrations occur at about twice, three-times, four-times and five-times the fundamental frequency. Peaks for C-H, N-H and O-H bonds are found in the NIR due to 2nd and 3rd overtones of these vibrations. In addition to the overtones, vibrations combine to form strong bands of absorbance, especially in the region from 5500~4000 cm-1. Thus, the NIR is divided into 3 regions, which roughly correspond to the combination region, and two overtone regions

## Slide 6: A Regra No 1
MUDANÇA QUÍMICA
=
MUDANÇA NO ESPECTRO NIR

## Slide 7: Interferômetro de Michelson
Espectro
Transformada de Fourier
Interferograma
Detector
Notes: FT technology makes it possible. IR source light (from a halogen lamp) is split into two beams. One beam bounces off a fixed mirror, the other bounces off a moving mirror. The waves of light interfere with each other to create an interferogram which is a change in intensity of light with time. This interferogram is transformed using a Fast Fourier Transform from the time domain to the frequency domain and a spectrum is produced.

## Slide 8: Interferômetro Yokogawa
Molas de fita
Sub VCM
Espêlho Parcial
Parte Fixa
Bobina VCM
Força de Atuação
Molas de fita
Parte Fixa
Notes: The Unique “Swing” Interferometer design makes the InfraSpec NR800 has no friction in the movement of the mirror, the mechanism never wears out.

## Slide 9: Interferograma e Espectro
Notes: Interferogram and Spectrum of Toluene.

## Slide 10: Fibra Óptica com Duplo Feixe
Ponto 1
Ponto 2
Ponto 3
Ponto 4
Processamento com DSP
  Aquisição de dados de 4 canais dos ADCs simultâneamente,
  Cálculo dos espectros em paralelo
DSP
Pontos de Amostragem
Fotodiodos InGaAs
ADC’s
Interferômetro
Divisores
CPU
Método Duplo Feixe
  Cancela efeitos do meio ambiente no índice de refração da fibra óptica
Notes: The InfraSpec NR800 is a reliable tool for the measurement of photoresists, washing solutions and other wet chemical treatments of wafers.
Individual component concentrations as well as physical properties can be measured and controlled in real time.
The NR800 is capable of up to 12 simultaneous measurements on 4 different channels, simultaneously.
The technology is non-destructive. Sample cells are usually placed in-line at circulation pumps, or can be probes inserted into tanks. No wastes are produced by the analyzer.
Available materials for cells and probes are virgin PTFE, electro-polished stainless steel, as well as exotic alloys.
Normal maintenance consists of changing the NIR source bulb every 6-7 months, and changing the laser about every 1.5 – 2 years.

## Slide 11: Configuração Típica para 2- Canais
Fibra Óptica
FT-NIR
DCS
Amostra
Retôrno ao processo
dreno
Amostra
Estação de trabalho NIR
Ethernet
dreno
Célula
Célula
Retôrno ao processo
RS-422 Modbus
Notes: One feature of the NR800 is the ability to support up to 4 simultaneous channels (cells or probes). A single analyzer can run samples at different points in the field, simultaneously. Also, each channel can measure up to 12 components or properties, simultaneously.
Other manufacturers use optical multiplexers, opto-mechanical devices that “clock” around from one set of fiber optic cable to another set. These devices are plagued by alignment problems and optical noise.

## Slide 12: Interferômetro do NR800
Estrutura anti-vibração

## Slide 13
Divisor óptico para configuração multi-canal
4 Canais de Fibra Óptica

## Slide 14
Célula tipo Flowthru
ou tipo Probe
Todos os Dados podem ser enviados via Ethernet, RS422 ou AO/DO para PLC, Datalogger/Controlador
Tipos de Sensor

## Slide 15
Condicionamento de Amostra

## Slide 16: Condicionamento de Amostra

## Slide 17: Yokogawa InfraSpec NR800
Aplicações nas Áreas de
Refino e Petroquímica

## Slide 18
O FT-NIR de Processo foi utilizado pela primeira vez em Refinarias no início dos anos 90.
Como as operações de Refino são realizadas com pequena margem de lucro, verificou-se, rapidamente, que o emprêgo do FT-NIR poderia:
Reduzir os Custos Operacionais
Garantir a Qualidade dos Produtos
Aplicações em Refino

## Slide 19: Aplicações na Área de Refino
RON
MON
CETANO
ÍNDICE CETANO
RVP
DENSIDADE
VISCOSIDADE
DESTILAÇÃO
PIONA
BENZENO
OXIGENADOS
MTBE
PONTO DE ENTUPIMENTO
PONTO DE FULGOR
PONTO DE CONGELAMENTO

## Slide 20: Espectro da Gasolina
Notes: This is a set of typical spectra of a regular gasoline. Note the changes in certain portions of the spectrum.

## Slide 21: 1a Derivada do Espectro
Notes: One of our first steps is to take the first derivative of the spectra. This eliminates baseline variations due to solids or bubbling of the sample, minor shifts in baseline from temperature changes in the analyzer, etc.

## Slide 22: Gasolina - RON
Notes: Then we correlate the active portions of the spectrum with the known concentrations. Here we show the calibration for RON. Note the range of analysis is 94 – 98 Octane Numbers and the accuracy (measured by SEC) is +/- 0.11% at the one sigma level.

## Slide 23: Gasolina - Densidade

## Slide 24: Gasolina - RVP
Notes: Here we have RVP in the 45 ~ 60 kPa range. Note the SEC of 0.65 kPa.
SEC => Standard Error of Calibration

## Slide 25: Gasolina - Benzeno
Notes: Here is the calibration of the Benzene in the range of 4 ~ 6% with an accuracy of 0.07% at the one sigma level.
Obviously, you can do this same analysis using a process GC. What you cannot do with a process GC is get 12 measured properties simultaneously in less than 1 minute.

## Slide 26: Gasolina - Aromáticos Totais
Notes: Total Aromatics 30 ~ 40% +/- 0.6% at one sigma.

## Slide 27: Gasolina Temperatura @ 50% Off
Notes: 50% Off Temperature, 80 ~ 100 oC +/- 1.2 oC

## Slide 28: Medição de Óleos pesados
Amostra
Parâmetro
Faixa Típica
RMSEC
Gasóleo
CFPP
-13 – 6 oC
1,75
PP ponto fluidez
-17 – 2 oC
1,99
Viscosidade
2,45 – 5,502 cSt
0,075
FP
55 – 109 oC
3,02
Peso Específico
0,8210 – 0,8400 g/cm3
0,0005
Óleo
Combustível
Enxôfre
0,0130 – 0,8590 %pêso
0,099
N Total
0,0205 – 0,02815 %pêso
0,023
Pêso Específico
0,86935 – 0,93975 g/cm3
0,00560
Notes: Cold Filter Plugging Point, Pour Point, Viscosity, Flash Point, Density, Total Nitrogen and Sulfur can also be determined in some samples. Table 1 below shows the typical ranges and REMSEC for measurements that have been performed.

## Slide 29: Temperatura do Óleo Combustível @ 50% Off

## Slide 30: Temperatura do Óleo Combustível @ 90% Off
Notes: Notice the scale of measurements is 340 ~ 375 oC

## Slide 31
O FT-NIR pose ser utilizado em inúmeras aplicações Petroquímicas.
Onde houver reatores, colunas de separação, misturas e outros processos, o NR800 da Yokogawa pode monitorar todos as correntes necessárias para tornar o processo mais eficiente.
Aplicações em Petroquímica

## Slide 32: Derivados Petroquímicos da Nafta e do Eteno
Etileno
Glicol
Etanolaminas
oxidação
Óxido de
Etileno
Nafta
Eteno
C2H4
Etanol
CH3-CH2OH
Acetaldeído
CH3-CH-O
Ácido Acético
CH3-COOH
Acetato Vinílico
H3C-C-O-CH=CH2
Éter Etílico
CH3-O-CH3
Hidrocarbonetos Halog.
CH3-CH2X
Refrigerantes
Di-Cloro-Etano
Cl-CH2-CH2-Cl
MVC
H2C=CH-Cl
Etil Benzeno
C6H5-CH2-CH3
Estireno
C6H5-CH-CH2

## Slide 33: Espectro NIR de amostra com diferentes teores de Estireno
Wave number(cm-1)
Absorbância(Abs)
Aromatic C-H stretch over tone
H
-C
2
5
C-H stretch over tone
=CH
2
C-H stretch over tone

## Slide 34: Medição de Estireno com NIR
80
81
82
83
84
85
80
81
82
83
84
85
Laboratório (%)
NIR (%)
Medição de Estireno
Faixa: 80,5～84,0%
SEP(σ): 0,089%
Correlação: 0,992

## Slide 35: Espectro NIR do Estireno com EB
Wave Number (cm-1)
Absorbância (Abs)

## Slide 36: Medição de Etil Benzeno com NIR
Laboratório (%)
83
84
85
83
84
85
NIR (%)
Etil Benzeno
Faixa: 83,0 ~ 84,8%
SEP (σ): 0,045%
Correlação: 0,996

## Slide 37: Derivados Petroquímicos do Benzeno, Tolueno, Xileno
Benzeno
Ácido Teraftálico
COOH
COOH
o, m, p-xilenos
CH3
CH3
Cumeno
Fenol
oxidação
Etil-Benzeno
Estireno
Ciclohexano
C6H12
OH
CHCH3-CH3
CH2-CH3
CH=CH2
Acetona
CH3-CO-CH3
Ácido Adípico
COOH-(CH2)4- COOH
Caprolactama
C6H9ONH
Nitrobenzeno
NO2
Anilina
NH2
Tolueno
CH3
Oxidação do p-xileno
BiFenol A
CH3CCH3
OH
OH

## Slide 38: Produção de m-Xileno
Faixa (%)
RMSEC
Wavenumber usado
Benzeno
0,424 ~ 1,383
0,018
8763 ~ 8300,
5492 ~ 5338
Tolueno
0,496 ~ 1,477
0,011
8855 ~ 8300, 5492 ~ 5338
Etil-Benzeno
3,74 ~ 23,37
0,066
8763 ~ 8300, 5492 ~ 5338
p-Xileno
4,06 ~ 22,08
0,095
8763 ~ 8300, 5492 ~ 5338
m-Xileno
4,51 ~ 51,06
0,174
8763 ~ 8300, 5492 ~ 5338

## Slide 39: Espectro do Xileno

## Slide 40: Benzeno em Xilenos
Notes: Note the range of 0.4 ~ 1.4% Benzene in Xylenes

## Slide 41: Medição de Água em Petroquímica
１．Em alguns casos a água atua como inibidor da reação
(Manutenção da atividade Catalítica)
２．Manter a qualidade do produto
（teor de água, padrão de qualidade, etc.）
３．Controle da Corrosão
（Corrosão por pitting em trocadores de calor)

## Slide 42
Laboratório (Método Karl Fisher)
  Requer pré-tratamento
  Êrros na amostragem (absorção de umidade)
  Requer Reagentes
FT-NIR (Rápido, Não destrutivo)
  On-line, tempo real
  Não ocorre absorção de umidade
  Não requer pré-tratamento
Teor de Água em Álcoois

## Slide 43
Espectro do Metanol com
diferentes teores de água
Teor de Água: 0~2000 ppm
Wave number(cm-1)
(Abs)
Espectro do Metanol
O-H stretch +
O-H def.
0 ppm
2000 ppm

## Slide 44: Lab. vs NIR
0
1000
2000
0
1000
2000
Laboratório
(ppm)
FT-NIR
(ppm)
Teor de Água no Metanol
Faixa: 0~2000ppm
SEP (σ): 26,9 ppm
Correlação: 0,999
Repetibilidade: 3 ppm

## Slide 45: Água em Óleo de Turbina

## Slide 46: Detalhe Óleo de Turbina

## Slide 47: 0~2000 ppm de Água em Óleo de Turbina
Notes: 0~2000 ppm +/-55 ppm Water in Turbine Oil

## Slide 48: Análise Multivariável On-Line de Propriedades e Composição
Analisadores Contínuos
  NDIR, UV, TC, Densidade, Umidade, pH, Cond., ORP
  Medida Específica
  Sistemas Binários
Cromatógrafos de Processo
  Para Amostras Gasosas e Líquidas que vaporizam de forma limpa
  Para Composição, mas não para Propriedades
Tituladores/Químico-in-a-Box
  Medida Específica (max. 2)
  Freqüente Manutenção – Reagentes – Resíduos - Mecânica

## Slide 49: Laboratório x NIR
Método Lab. Método NIR
Contínuo Não Sim
On-line Não Sim
Multi-componente Não Sim
Não-destrutiva Não Sim
Reagentes, resíduos Sim Não
Manutenção Alta Baixa

## Slide 50: Vantagens do NR800
Tecnologia FT-NIR – informação espectral precisa
Interferômetro exclusivo - alto S/N, resistente à vibração e mudanças de temperatura
Conexão ao processo por Fibra Óptica - segurança
Fibra Óptica de Duplo Feixe - linha base estável
Multi-canal simultâneo - sem uso de multiplexadores
Computador On-board - conceito distribuído
Modelo de calibração robusto - fácil de calibrar
Rápido, Direto e Contínuo (10min PGC x 30seg NIR)
Substitui análises caras e demoradas no Laboratório por medições automáticas e em tempo real
Análise Não-destrutiva, Sem reagentes, Sem pré-tratamento
Medição Multi-componente/Multi-propriedade, Multivariável
Notes: The main advantages of the NR800 compared to other process FTNIR’s

## Slide 51: NIR x Mid-IR
I = Ioe-KCL
K: coef. Absorbância
C: Concentração
L: Path length
Io
Luz
L
Lei de Beer-Lambert
Sensor
Io
Amostra
I
Near IR
Fraca1/100 - 1/1000
(longer cell paths)
Picos largos
Ricos em Informação
Líquido 1mm para Água
SIM
SIM
SIM
SIM
Modelo Regressão
( Chemometrics )
Mid IR
Forte
(short cell paths)
Picos estreitos
Específicos
Líquido 0,01mm Gas 1mm
NÃO
SIM
SIM
SIM
Absoluta
Absorbância
Espectro
Path Length
Fibra Óptica
(Sensor Remoto)
Medição de Propriedades
Medição de
Multicomponentes
Medição
Contínua
Calibração
Notes: Mid-IR has certain features that in some cases are of advantage, but most often are disadvantages for process measurements:
1.) Because of the strong Mid-IR absorbance, cell path lengths for liquid samples are ridiculously short. It is difficult to build and flush such small cells due to their high pressure drop.
2.)The broader, more complex bands of absorbance in the NIR allows us to measure more complex properties such as density, octane, average molecular weight and other properties that cannot be calculated from Mid-IR spectra.
3.) The optical components required by Mid-IR are often hydroscopic and fragile.
Note the one thing we don’t do on the NR800: gases.

## Slide 52: Amostras para Modelamento
Regras
  1. No. de amostras necessárias: 30~50
  2. Volume das amostras:
aprox. 300 ml para instalação em Lab.
aprox. 1 litro para instalação On-line
  3. Recipiente: Duplo selo, em Vidro
  4. “Boa” análise Laboratorial

## Slide 53: Coleta da Amostra
As propriedades reais das amostras devem estar bem distribuídas em toda a faixa requerida do analisador. (Colha amostras bem distribuídas)
Pode ser necessária a coleta de amostras durante um longo período para obtenção de amostras com faixas suficientemente amplas para calibração e validação.
Identifique cuidadosamene os Dados do Lab com os Espectros NIR.
Evite misturas manuais das amostras que não representam o processo real.